## Cover
DO NOT SAVE AS A GOOGLE SHEET, GOOGLE STRIPS OUT CONDITIONAL FORMATTING WHICH HIGHLIGHTS THE BEST REFERENCE STANDARD IN CELLS A29:D31
Spreadsheet to determine the best reference |  |  |  | Reference Standards List
standard to perform a calibration |  |  |  | PRESSURE | RANGES | ACCURACY | TUR | Controller | Note
|  |  |  | Fluke 2465/8A | 15-0 PSI | .0018% RD + .00000145 PSI | .0018% RD+.01 Pa | Yes | .01 Pa = .00000145 PSI
|  |  |  | Mensor CPC2000 | .4" | .3% FS |  | Yes | Can monitor with MC6 for higher accuracy
Device under test | Enter below * |  |  | MC6 Module | +/-0.125" | .07% FS | .00021"H2O
Device under test range |  |  |  | MC6 Module | +/-.25" | .07% FS | .00038"H2O
Device under test % of reading accuracy |  | B7 or B8 must be 0 |  | MC6 Module | .5" | .07% FS | .00038"H2O
Device under test % of full scale accuracy |  | B7 or B8 must be 0 |  | MC6 Module | 1" | .06% FS | .0006"H2O
|  |  |  | MC6 Module | +/-2.5" | .06% FS | .0031"H2O
Mixed % RD + % FS standard |  |  |  | MC6 Module | +/-5" | .06% FS | .0062"H2O
PACE6000 module range |  |  |  | MC6 Module | 10" | .06% FS | .0062"H2O
PACE6000 % of reading accuracy |  |  |  | MC6 Module | +/-12.5" | .06% FS | .016"H2O
PACE6000 % of full scale accuracy |  |  |  | MC6 Module | +/-25" | .06% FS | .038"H2O
|  |  |  | MC6 Module | 1 - 200" | .06% FS typical, some .1% FS
% full scale standard |  |  |  | PACE | (2) 10" | .05% RD + .05% FS |  | Yes
MC6 module or Mensor range |  |  |  | Ruska 7250LP | Dual 10" / 30" | 3-30" .009% RD, 0-3" .009% FS (.00027") | .0068% RD+.00076 "H2O | Yes
MC6 module or Mensor accuracy |  |  |  | Minerva 700-VLP | 1.6" / 8" / 100" | 8-100" .01% RD, 1.6-8" .02% RD, 0-1.6" .0004" |  | Yes
|  |  |  | Druck DPI-515 | 1 PSI | .03% FS |  | Yes
% of reading standard |  |  |  | PACE | 1 PSI | .025% RD + .025% FS |  | Yes
Deadweight tester or Ruska range |  |  |  | MC6 Module | 5 - 300 PSI | .05% FS typical, some .025% and .1% FS
Deadweight tester or Ruska accuracy |  |  |  | PACE | 5 PSI | .02% RD + .02% FS |  | Yes
|  |  |  | PACE | 10 PSI | .005% RD + .005% FS |  | Yes
* values must be in the same units for all fields, DO NOT type W.C., PSI, or % in any field |  |  |  | PACE | 15 PSI | .02% RD + .02% FS, R54 has a .005% + .005% |  | Yes | CM0 is low accuracy CM2 is high accuracy
|  |  |  | Fluke 2465/8A | 0-50 PSI | .0018% RD + .00000145 PSI | .0018% RD+.01 Pa | Yes | .01 Pa = .00000145 PSI
|  |  |  | PACE | 100 PSI | .005% RD + .005% FS |  | Yes
The best standard to use is highlighted |  |  | Test Accuracy Ratio | Druck DPI-515 | 100 PSI | .01% FS |  | Yes
|  |  | Must be greater than 4** | PACE | 7 BAR / 101 PSI | .02% RD + .02% FS |  | Yes
Mixed % RD + %FS standard - Pace6000 |  |  | 0 | PACE | 150 PSI | .005% RD + .005% FS |  | Yes
% of full scale standard - MC6, Mensor 15000, Druck |  |  | 0 | PACE | 10 BAR / 145 PSI | .005% RD + .005% FS |  | Yes
% of reading standard - Deadweight tester, Ruska 7250LP |  |  | 0 | PACE | 290 PSI | .005% RD + .005% FS |  | Yes
** If none of the master gage options provide a 4 to 1 test ratio or greater review with Robert |  |  |  | MC6 Module | 300 PSI | .05% FS |  |  | Should use 300 PSI Pace or Mensor instead
|  |  |  | Mensor 15000 | 300 PSI | .01% FS
|  |  |  | PACE | 300 PSI | .005% RD + .005% FS |  | Yes
|  |  |  | PACE | 300 PSI + ABS | .005% RD + .005% FS |  | Yes | Barometric sensor failed CAN'T USE FOR CERTS
|  |  |  | Deadweight | 3 - 500 PSI | 50-500 PSI .015% RD, 0-50 .015% FS (.0075 PSI) |  | Yes | .008% RD only when using the Fluke software
|  |  |  | PACE | 500 PSI + ABS | .005% RD + .005% FS |  | Yes | Has barometric sensor for absolute readings
|  |  |  | MC6 Module | 600 PSI | .05% FS
|  |  |  | MC6 Module | 1000 PSI | .025% and .05% FS |  |  | Should use 1000 PSI Pace or Mensor instead
|  |  |  | Mensor CPC3000 | 1000 PSI | .025% FS |  | Yes | Should use 1000 PSI Pace or Mensor instead
|  |  |  | Mensor 15000 | 1000 PSI | .01% FS |  |  | Should use 1000 PSI Pace or Mensor instead
|  |  |  | PACE | (2) 1000 PSI | .005% RD + .005% FS |  | Yes
|  |  |  | Fluke 2465/8A | 50-1000 PSI | .0033% RD + .00010 PSI | .0033% RD+.27 Pa | Yes | .71 Pa = .00010 PSI
|  |  |  | PACE | 70 BAR / 1015 PSI | .005% RD + .005% FS |  | Yes
|  |  |  | PACE | 3000 PSI | .005% RD + .005% FS |  | Yes
|  |  |  | MC6 Module | 7500 PSI | .025% FS
|  |  |  | Mensor 15000 | 10000 PSI | .02% FS |  |  | Derated from .015% last time it was certified
VELOCITY
|  |  |  | TSI | 200 - 750 FPM | 2% RD + .5% FS
|  |  |  | TSI | 1000 - 5000 FPM | 2% RD + .5% FS
|  |  |  | TSI | 7500 and 10000 FPM | 2% RD + .5% FS
|  |  |  | EE75 | 40 - 8000 FPM | 1% RD + .5% FS
|  |  |  | TEMPERATURE |  | TEMPERATURE SOURCES
|  |  |  | Azonix | -328 +/- .4F | Glycol bath -20 to +74C (-4 to 162F)
|  |  |  | Azonix | 32 +/- .09F | Hart 9122 drywell 35 to 600C (95 to 1700F)
|  |  |  | Azonix | 212 +/- .4F | Thermotrons big -70 to 170C (-100 to 350F)
|  |  |  | Azonix | 392 +/- .7F | Thermotrons small -50 to 190C (-65 to 375F)
Espec low temp -85 to 180C (-120 to 354F)
|  |  |  | Hart 1502A w/5616 | -323 +/- .029F | Espec high temp 20 to 300C (70 to 570F)
|  |  |  | Hart 1502A w/5616 | 32 +/- .031F
|  |  |  | Hart 1502A w/5616 | 392 +/- .043F
|  |  |  | Hart 1502A w/5616 | 788 +/- .083F
HUMIDITY
|  |  |  | Thunder Scientific | 10 - 95% | .5% RH
|  |  |  | Thunder Scientific | 0 - 70C, 32 - 158F | Limit maximum for certs at 40 - 130F
|  |  |  | General Eastern | 10 - 95% | 1.2% RH
|  |  |  | Viasala PTU300 | 0 - 90% | 1.0% RH @ 72 Deg. F
|  |  |  | Viasala PTU300 | 90 - 100% | 1.7% RH @ 72 Deg. F
|  |  |  | Viasala HMT330 | 0 - 90% | 1.0% RH @ 72 Deg. F
|  |  |  | Viasala HMT330 | 90 - 100% | 1.7% RH @ 72 Deg. F

## Calculations
|  |  | Pace6000 | Accuracy %RD + %FS |  |  |  | MC6 or % full scale standard |  |  |  |  | Deadweight tester or % reading standard
|  |  | Pace module range |  |  |  |  | Type in range |  |  |  |  | Type in range
|  |  | % of reading error |  |  |  |  | % full scale error |  |  |  |  | % of reading error
% full scale error
|  |  | Pressure | Max low | Max high | Total % RD+FS error |  | Pressure | Max low | Max high | % FS error |  | Pressure | Max low | Max high | % RD error
|  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  |  | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE! | #VALUE!
Range of DUT
% RD error of DUT
% FS error of DUT
Pressure | Max low | DUT allowable % error | Pressure | Max low | Max high | Total RD+FS error | Ratio at point | Pressure | Max low | Max high | % FS error | Ratio at point | Pressure | Max low | Max high | % RD error | Ratio at point
#VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
#VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
#VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
#VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
#VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
#VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
#VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
#VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
#VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| Worst case point | #VALUE! |  |  | Worst case point | 100 | 0 |  |  | Worst case point | 100 | 0 |  |  | Worst case point | 100 | 0
1
1
1

## Equipment error comparison
Pressure | Error at a specific test point
"W.C. | Ruska 7250LP | Ashcroft .25" | Ashcroft .5" | Ashcroft 1" | Ashcroft 5" | Ashcroft 10" | PACE 10" | PACE 1 PSI | Ashcroft 100" | Ashcroft 5 PSI | PACE 5 PSI | Minerva | Equipment | Accuracy
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ruska 7250LP | 0-3" .009%FS, 3-30" .009%RD
0.01 | 0.00027 | 0.0175 | 0.035 | 0.06 | 0.3 | 0.6 | 0.5005 | 0.693 | 6 | 3.4635 | 2.771 | 0.0004 | Ashcroft .25" | .07%FS
0.02 | 0.00027 | 0.00875 | 0.0175 | 0.03 | 0.15 | 0.3 | 0.2505 | 0.34663 | 3 | 1.7317 | 1.3856 | 0.0004 | Ashcroft .5" | .07%FS
0.03 | 0.00027 | 0.0058333 | 0.011667 | 0.02 | 0.1 | 0.2 | 0.16717 | 0.23117 | 2 | 1.1545 | 0.9238 | 0.0004 | Ashcroft 1" | .06%FS
0.04 | 0.00027 | 0.004375 | 0.00875 | 0.015 | 0.075 | 0.15 | 0.1255 | 0.17344 | 1.5 | 0.86587 | 0.6929 | 0.0004 | Ashcroft 5" | .06%FS
0.05 | 0.00027 | 0.0035 | 0.007 | 0.012 | 0.06 | 0.12 | 0.1005 | 0.1388 | 1.2 | 0.6927 | 0.55436 | 0.0004 | Ashcroft 10" | .06%FS
0.06 | 0.00027 | 0.0029167 | 0.0058333 | 0.01 | 0.05 | 0.1 | 0.083833 | 0.11571 | 1 | 0.57725 | 0.462 | 0.0004 | PACE 10" | .05%RD + .05%FS
0.07 | 0.00027 | 0.0025 | 0.005 | 0.0085714 | 0.042857 | 0.085714 | 0.071929 | 0.099214 | 0.85714 | 0.49479 | 0.39603 | 0.0004 | PACE 1 PSI | .025% RD + .025% FS
0.08 | 0.00027 | 0.0021875 | 0.004375 | 0.0075 | 0.0375 | 0.075 | 0.063 | 0.086844 | 0.75 | 0.43294 | 0.34655 | 0.0004 | Ashcroft 100" | .06%FS
0.09 | 0.00027 | 0.0019444 | 0.0038889 | 0.0066667 | 0.033333 | 0.066667 | 0.056056 | 0.077222 | 0.66667 | 0.38483 | 0.30807 | 0.0004 | Ashcroft 5 PSI | .025%FS
0.1 | 0.00027 | 0.00175 | 0.0035 | 0.006 | 0.03 | 0.06 | 0.0505 | 0.069525 | 0.6 | 0.34635 | 0.27728 | 0.0004 | PACE 5 PSI | .02% RD + .02% FS
0.2 | 0.00027 | 0.000875 | 0.00175 | 0.003 | 0.015 | 0.03 | 0.0255 | 0.034888 | 0.3 | 0.17317 | 0.13874 | 0.0004 | Minerva | 0-1.6" .0004", 1.6-8" .02%RD, 8-100" .01%RD
0.3 | 0.00027 |  | 0.0011667 | 0.002 | 0.01 | 0.02 | 0.017167 | 0.023342 | 0.2 | 0.11545 | 0.09256 | 0.0004
0.4 | 0.00027 |  | 0.000875 | 0.0015 | 0.0075 | 0.015 | 0.013 | 0.017569 | 0.15 | 0.086587 | 0.06947 | 0.0004
0.5 | 0.00027 |  | 0.0007 | 0.0012 | 0.006 | 0.012 | 0.0105 | 0.014105 | 0.12 | 0.06927 | 0.055616 | 0.0004
0.6 | 0.00027 |  |  | 0.001 | 0.005 | 0.01 | 0.0088333 | 0.011796 | 0.1 | 0.057725 | 0.04638 | 0.0004
0.7 | 0.00027 |  |  | 0.00085714 | 0.0042857 | 0.0085714 | 0.0076429 | 0.010146 | 0.085714 | 0.049479 | 0.039783 | 0.0004
0.8 | 0.00027 |  |  | 0.00075 | 0.00375 | 0.0075 | 0.00675 | 0.0089094 | 0.075 | 0.043294 | 0.034835 | 0.0004
0.9 | 0.00027 |  |  | 0.00066667 | 0.0033333 | 0.0066667 | 0.0060556 | 0.0079472 | 0.066667 | 0.038483 | 0.030987 | 0.0004
1 | 0.00027 |  |  | 0.0006 | 0.003 | 0.006 | 0.0055 | 0.0071775 | 0.06 | 0.034635 | 0.027908 | 0.0004
2 | 0.00027 |  |  |  | 0.0015 | 0.003 | 0.003 | 0.0037137 | 0.03 | 0.017317 | 0.014054 | 0.002
3 | 9e-05 |  |  |  | 0.001 | 0.002 | 0.0021667 | 0.0025592 | 0.02 | 0.011545 | 0.009436 | 0.002
4 | 9e-05 |  |  |  | 0.00075 | 0.0015 | 0.00175 | 0.0019819 | 0.015 | 0.0086588 | 0.007127 | 0.002
5 | 9e-05 |  |  |  | 0.0006 | 0.0012 | 0.0015 | 0.0016355 | 0.012 | 0.006927 | 0.0057416 | 0.002
6 | 9e-05 |  |  |  |  | 0.001 | 0.0013333 | 0.0014046 | 0.01 | 0.0057725 | 0.004818 | 0.002
7 | 9e-05 |  |  |  |  | 0.00085714 | 0.0012143 | 0.0012396 | 0.0085714 | 0.0049479 | 0.0041583 | 0.002
8 | 9e-05 |  |  |  |  | 0.00075 | 0.001125 | 0.0011159 | 0.0075 | 0.0043294 | 0.0036635 | 0.001
9 | 9e-05 |  |  |  |  | 0.00066667 | 0.0010556 | 0.0010197 | 0.0066667 | 0.0038483 | 0.0032787 | 0.001
10 | 9e-05 |  |  |  |  | 0.0006 | 0.001 | 0.00094275 | 0.006 | 0.0034635 | 0.0029708 | 0.001
20 | 9e-05 | [6 empty] | 0.00059638 | 0.003 | 0.0017318 | 0.0015854 | 0.001
30 | 9e-05 | [7 empty] | 0.002 | 0.0011545 | 0.0011236 | 0.001
40 | [8 empty] | 0.0015 | 0.00086588 | 0.0008927 | 0.001
50 | [8 empty] | 0.0012 | 0.0006927 | 0.00075416 | 0.001
60 | [8 empty] | 0.001 | 0.00057725 | 0.0006618 | 0.001
70 | [8 empty] | 0.00085714 | 0.00049479 | 0.00059583 | 0.001
80 | [8 empty] | 0.00075 | 0.00043294 | 0.00054635 | 0.001
90 | [8 empty] | 0.00066667 | 0.00038483 | 0.00050787 | 0.001
100 | [8 empty] | 0.0006 | 0.00034635 | 0.00047708 | 0.001
Pressure | Error at a specific test point
"W.C. | PACE 15 PSI | PACE 100 PSI | Ashcroft 100 PSI | PACE 300 PSI | Fluke Deadweight | PACE 500 PSI |  | Equipment |  | Accuracy
0 | 0 | 0 | 0 | 0 | 0 | 0 |  | PACE 15 PSI |  | .02% RD + .02% FS
5 | 0.0008 | 0.00105 | 0.012 | 0.00305 | 0.0075 | 0.00505 |  | PACE 100 PSI |  | .005%RD + .005%FS
10 | 0.0005 | 0.00055 | 0.006 | 0.00155 | 0.0075 | 0.00255 |  | Ashcroft 100 PSI |  | .06%FS
15 | 0.0004 | 0.00038333 | 0.004 | 0.00105 | 0.0075 | 0.0017167 |  | PACE 300PSI |  | .005%RD + .005%FS
20 |  | 0.0003 | 0.003 | 0.0008 | 0.0075 | 0.0013 |  | Fluke Deadweight |  | 50-500 PSI .015% RD, 0-50 .015% FS (.0075 PSI)
25 |  | 0.00025 | 0.0024 | 0.00065 | 0.0075 | 0.00105 |  | PACE 500PSI |  | .005%RD + .005%FS
30 |  | 0.00021667 | 0.002 | 0.00055 | 0.0075 | 0.00088333
35 |  | 0.00019286 | 0.0017143 | 0.00047857 | 0.0075 | 0.00076429
40 |  | 0.000175 | 0.0015 | 0.000425 | 0.0075 | 0.000675
45 |  | 0.00016111 | 0.0013333 | 0.00038333 | 0.0075 | 0.00060556
50 |  | 0.00015 | 0.0012 | 0.00035 | 0.0075 | 0.00055
100 |  | 0.0001 | 0.0006 | 0.0002 | 0.00015 | 0.0003
150 |  |  |  | 0.00015 | 0.00015 | 0.00021667
200 |  |  |  | 0.000125 | 0.00015 | 0.000175
250 |  |  |  | 0.00011 | 0.00015 | 0.00015
300 |  |  |  | 0.0001 | 0.00015 | 0.00013333
350 |  |  |  |  | 0.00015 | 0.00012143
400 |  |  |  |  | 0.00015 | 0.0001125
450 |  |  |  |  | 0.00015 | 0.00010556
500 |  |  |  |  | 0.00015 | 0.0001