## Revision Log
REVISION LOG: Data Template - 34401A
Revision Number | Date | Comments | Author | Function
0 | 2020-05-12 | Initial creation of spreadsheet. | Robert A. Maldonado | Associate Metrology Engineer
1 | 2020-10-23 | Updated format of spreadsheet to 1) print correctly, 2) have gridlines and borders, 3) have updated fields / title / headers, 4) additional information relating to customer, product and calibration process information and 5) include updated tolerances based on guard banded uncertainties - best case along with 6) entry fields for reference data. 50% tolerance calculation is removed. Percent deviation is provided so technicians and quality personnel can review percentage of tolerance used by product and determine of adjustment is necessary based on internal measurement assurance program rules. | Robert A. Maldonado | Associate Metrology Engineer

## Data Template
Data Template - 34401A
Customer, Product and Calibration Process Information
Certificate Number: |  |  |  | Temperature (F): |  | Procedure 1: |  | TC-00062
Sales Order Number: |  |  |  | Humidity (%RH): |  | Procedure 2: |  | TC-00072
Customer P.O. #: |  |  |  | Pressure (inHg): |  | Procedure 3: |  | TC-00073
Date of Receipt: |  |  |  | Dewpoint (F): |  | Procedure 4: |  | TC-00074
Calibration Date: |  |  |  | UUT Serial Number:
Calibration Due Date: |  |  |  | UUT Customer ID:
Zero Offset Functions (No comparison to reference standard; Only data obtained from UUT)
DC Current (Open terminals)
|  | As Found |  |  |  | As Left
Zero Offset | Tolerance | Nominal | 34401A | Percent Error | Pass / Fail | Nominal | 34401A | Percent Error | Pass / Fail
10mA Range | 0.002 | 0 |  |  |  | 0
100mA Range | 0.005 | 0 |  |  |  | 0
1A Range | 0.0001 | 0 |  |  |  | 0
3A Range | 0.0006 | 0 |  |  |  | 0
DC Voltage (Shorted terminals)
|  | As Found |  |  |  | As Left
Zero Offset | Tolerance | Nominal | 34401A | Percent Error | Pass / Fail | Nominal | 34401A | Percent Error | Pass / Fail
100mV Range | 0.0035 | 0 |  |  |  | 0
1V Range | 7e-06 | 0 |  |  |  | 0
10V Range | 5e-05 | 0 |  |  |  | 0
100V Range | 0.0006 | 0 |  |  |  | 0
1000V Range | 0.01 | 0 |  |  |  | 0
Resistance 4-Wire (Shorted terminals)
|  | As Found |  |  |  | As Left
Zero Offset | Tolerance | Nominal | 34401A | Percent Error | Pass / Fail | Nominal | 34401A | Percent Error | Pass / Fail
100 Ohms Range | 0.004 | 0 |  |  |  | 0
1k Ohms Range | 1e-05 | 0 |  |  |  | 0
10k Ohms Range | 0.0001 | 0 |  |  |  | 0
100k Ohms Range | 0.001 | 0 |  |  |  | 0
1M Ohms Range | 1e-05 | 0 |  |  |  | 0
10M Ohms Range | 0.0001 | 0 |  |  |  | 0
100M Ohms Range | 0.01 | 0 |  |  |  | 0
Gain Verification
DC Voltage
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34401A | Percent Error | Pass / Fail | Reference | 34401A | Percent Error | Pass / Fail
100mV | 0.0085
1V | 4.7e-05
10V | 0.00037
-10V | 0.00037
100V | 0.0051
1000V | 0.055
Data Template - 34401A
Gain Verification (Continued)
AC Voltage (UUT AC Filter must be set to slow mode - 3 Hz)
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34401A | Percent Error | Pass / Fail | Reference | 34401A | Percent Error | Pass / Fail
10mV @1kHz | 0.046
100mV @1kHz | 0.1
100mV @50kHz | 0.17
1V @1kHz | 0.0009
1V @50kHz | 0.0017
10V @ 10Hz | 0.038
10V @1kHz | 0.017
10V @50kHz | 0.017
100V @1kHz | 0.09
100V @50kHz | 0.17
750V @1kHz | 0.675
4 Wire Resistance
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34401A | Percent Error | Pass / Fail | Reference | 34401A | Percent Error | Pass / Fail
100 Ohms | 0.014
1k Ohms | 0.00011
10k Ohms | 0.0011
2 Wire Resistance
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34401A | Percent Error | Pass / Fail | Reference | 34401A | Percent Error | Pass / Fail
100k Ohms | 0.0104
1M Ohms | 0.0001
10M Ohms | 0.0041
100M Ohms | 0.81
DC Current
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34401A | Percent Error | Pass / Fail | Reference | 34401A | Percent Error | Pass / Fail
10mA | 0.007
100mA | 0.0538
1A | 0.000974
2A | 0.00287
AC Current (AC filter slow)
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34401A | Percent Error | Pass / Fail | Reference | 34401A | Percent Error | Pass / Fail
1A @ 1kHz | 0.0014
2A @ 1kHz | 0.0048
Frequency
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34401A | Percent Error | Pass / Fail | Reference | 34401A | Percent Error | Pass / Fail
100Hz @ 10mV | 0.01
100kHz @ 1V | 0.01