## Revision Log
REVISION LOG: Data Template - 34450A
Revision Number | Date | Comments | Author | Function
0 | 2020-05-12 | Initial creation of spreadsheet. | Robert A. Maldonado | Associate Metrology Engineer
1 | 2020-10-23 | Updated format of spreadsheet to 1) print correctly, 2) have gridlines and borders, 3) have updated fields / title / headers, 4) additional information relating to customer, product and calibration process information and 5) include updated tolerances based on guard banded uncertainties - best case along with 6) entry fields for reference data. 50% tolerance calculation is removed. Percent deviation is provided so technicians and quality personnel can review percentage of tolerance used by product and determine of adjustment is necessary based on internal measurement assurance program rules. | Robert A. Maldonado | Associate Metrology Engineer

## Data Template
Data Template - 34450A
Customer, Product and Calibration Process Information
Certificate Number: |  |  |  | Temperature (F): |  | Procedure 1: |  | TC-00062
Sales Order Number: |  |  |  | Humidity (%RH): |  | Procedure 2: |  | TC-00072
Customer P.O. #: |  |  |  | Pressure (inHg): |  | Procedure 3: |  | TC-00074
Date of Receipt: |  |  |  | Dewpoint (F): |  | Procedure 4: |  | TC-00075
Calibration Date: |  |  |  | UUT Serial Number:
Calibration Due Date: |  |  |  | UUT Customer ID:
Zero Offset Functions (No comparison to reference standard; Only data obtained from UUT)
DC Current (Open terminals - 5 1/2 Digit mode)
|  | As Found |  |  |  | As Left
Zero Offset | Tolerance | Nominal | 34450A | Percent Error | Pass / Fail | Nominal | 34450A | Percent Error | Pass / Fail
100uA | 0.015 | 0 |  |  |  | 0
1mA | 7e-05 | 0 |  |  |  | 0
10mA | 0.0015 | 0 |  |  |  | 0
100mA | 0.007 | 0 |  |  |  | 0
1A | 0.00015 | 0 |  |  |  | 0
10A | 0.0007 | 0 |  |  |  | 0
Capacitance (Open terminals - 5 1/2 Digit mode)
|  | As Found |  |  |  | As Left
Zero Offset | Tolerance | Nominal | 34450A | Percent Error | Pass / Fail | Nominal | 34450A | Percent Error | Pass / Fail
1nF | No Specification provided; Test to verify functionality but do not record zero offset value at 1 nF
10nF | 0.05 | 0 |  |  |  | 0
100nF | 0.5 | 0 |  |  |  | 0
1uF | 0.005 | 0 |  |  |  | 0
10uF | 0.05 | 0 |  |  |  | 0
100uF | 0.5 | 0 |  |  |  | 0
1mF | 0.005 | 0 |  |  |  | 0
10mF | 0.05 | 0 |  |  |  | 0
DC Voltage (Shorted terminals - 5 1/2 Digit mode)
|  | As Found |  |  |  | As Left
Zero Offset | Tolerance | Nominal | 34450A | Percent Error | Pass / Fail | Nominal | 34450A | Percent Error | Pass / Fail
100mV | 0.008 | 0 |  |  |  | 0
1V | 5e-05 | 0 |  |  |  | 0
10V | 0.0005 | 0 |  |  |  | 0
100V | 0.005 | 0 |  |  |  | 0
1kV | 0.05 | 0 |  |  |  | 0
Resistance 2-Wire (Shorted terminals - 5 1/2 Digit mode)
|  | As Found |  |  |  | As Left
Zero Offset | Tolerance | Nominal | 34450A | Percent Error | Pass / Fail | Nominal | 34450A | Percent Error | Pass / Fail
100 Ohms | 0.208 | 0 |  |  |  | 0
1k Ohms | 0.00028 | 0 |  |  |  | 0
10k Ohms | 0.0007 | 0 |  |  |  | 0
100k Ohms | 0.005 | 0 |  |  |  | 0
1M Ohms | 5e-05 | 0 |  |  |  | 0
10M Ohms | 0.0005 | 0 |  |  |  | 0
100M Ohms | 0.005 | 0 |  |  |  | 0
Resistance 4-Wire (Shorted terminals - 5 1/2 Digit mode)
|  | As Found |  |  |  | As Left
Zero Offset | Tolerance | Nominal | 34450A | Percent Error | Pass / Fail | Nominal | 34450A | Percent Error | Pass / Fail
100 Ohms | 0.008 | 0 |  |  |  | 0
1k Ohms | 8e-05 | 0 |  |  |  | 0
10k Ohms | 0.0005 | 0 |  |  |  | 0
100k Ohms | 0.005 | 0 |  |  |  | 0
1M Ohms | 5e-05 | 0 |  |  |  | 0
10M Ohms | 0.0005 | 0 |  |  |  | 0
100M Ohms | 0.005 | 0 |  |  |  | 0
Data Template - 34450A
Gain Verification
DC Voltage
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34450A | Percent Error | Pass / Fail | Reference | 34450A | Percent Error | Pass / Fail
100mV | 0.026
-100mV | 0.026
1V | 0.0002
-1V | 0.0002
10V | 0.002
100V | 0.02
1kV | 0.2
AC Voltage
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34450A | Percent Error | Pass / Fail | Reference | 34450A | Percent Error | Pass / Fail
100mV @20Hz | 1.1
100mV @45Hz | 0.3
100mV @1kHz | 0.3
100mV @10kHz | 0.3
100mV @30kHz | 1.8
100mV @100kHz | 3.3
1V @20Hz | 0.011
1V @45Hz | 0.003
1V @1kHz | 0.003
1V @10kHz | 0.003
1V @30kHz | 0.018
1V @100kHz | 0.033
10V @ 20Hz | 0.11
10V @ 45Hz | 0.03
10V @ 1KkHz | 0.03
10V @ 10kHz | 0.03
10V @ 30kHz | 0.18
10V @ 100kHz | 0.33
100V @45Hz | 0.3
100V @1kHz | 0.3
100V @10kHz | 0.3
100V @30kHz | 1.8
100V @100kHz | 3.3
750V @45Hz | 2.25
750V @1kHz | 2.25
750V @10kHz | 2.25
Capacitance
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34450A | Percent Error | Pass / Fail | Reference | 34450A | Percent Error | Pass / Fail
1nF | No Specification provided; Test to verify functionality but do not record zero offset value at 1 nF
10nF | 0.15
100nF | 1.5
1uF | 0.015
10uF | 0.15
100uF | 1.5
1mF | 0.015
10mF | 0.25
Data Template - 34450A
Gain Verification (Continued)
Resistance 2-Wire
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34450A | Percent Error | Pass / Fail | Reference | 34450A | Percent Error | Pass / Fail
100 Ohms | 0.058
1k Ohms | 0.00058
10k Ohms | 0.0055
100k Ohms | 0.055
1M Ohms | 0.00065
10M Ohms | 0.0255
100M Ohms | 2.005
Resistance 4-Wire
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34450A | Percent Error | Pass / Fail | Reference | 34450A | Percent Error | Pass / Fail
100 Ohms | 0.058
1k Ohms | 0.00058
10k Ohms | 0.0055
100k Ohms | 0.055
1M Ohms | 0.00065
10M Ohms | 0.0255
100M Ohms | 2.005
DC Current
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34450A | Percent Error | Pass / Fail | Reference | 34450A | Percent Error | Pass / Fail
100uA | 0.065
1mA | 0.00057
10mA | 0.0065
100mA | 0.057
1A | 0.00115
10A | 0.0257
AC Current
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34450A | Percent Error | Pass / Fail | Reference | 34450A | Percent Error | Pass / Fail
10mA @ 20Hz | 0.16
10mA @ 45Hz | 0.06
10mA @ 1kHz | 0.06
10mA @ 5kHz | 0.22
10mA @ 10kHz | 0.22
100mA @ 20Hz | 1.6
100mA @ 45Hz | 0.6
100mA @ 1kHz | 0.6
100mA @ 5kHz | 2.2
100mA @ 10kHz | 2.2
1A @ 20Hz | 0.016
1A @ 45Hz | 0.006
1A @ 1kHz | 0.006
1A @ 5kHz | 0.022
2A @ 5kHz | 0.06
10A @ 45Hz | 0.06
10A @ 1kHz | 0.06
Frequency
|  | As Found |  |  |  | As Left
Set Point | Tolerance | Reference | 34450A | Percent Error | Pass / Fail | Reference | 34450A | Percent Error | Pass / Fail
20Hz @ .1V | 0.007
1kHz @ 1V | 0.00023