## Revision Log
REVISION LOG: Data Template - 34450A
Revision Number | Date | Comments | Author | Function
0 | 2020-05-12 | Initial creation of spreadsheet. | Robert A. Maldonado | Associate Metrology Engineer

## Data Template
Data Template - 34450A
Plant | R64 | TEMP - HUM - PRESS - DEW |  |  | 73.2 40.1 29.5 47.5
Serial Number | MY57092235 | MW-DMM | 15 | XXDWY00- |  | Date | 2020-05-11
|  | As Found |  |  | As Left
Zero Offset | Tolerance | 34450A | 100% | 50% | 34450A | 100% | 50%
DC Current (Open terminals - 5 1/2 Digit mode)
100uA | 0.015 | -0.006 | PASS | PASS | -0.006 | PASS | PASS
1mA | 7e-05 | 1e-05 | PASS | PASS | 1e-05 | PASS | PASS
10mA | 0.0015 | 0.0006 | PASS | PASS | 0.0006 | PASS | PASS
100mA | 0.007 | 0.001 | PASS | PASS | 0.001 | PASS | PASS
1A | 0.00015 | -7e-05 | PASS | PASS | -7e-05 | PASS | PASS
10A | 0.0007 | -0.0001 | PASS | PASS | -0.0001 | PASS | PASS
Capacitance (Open terminals - 5 1/2 Digit mode)
1nF | No Specification, skip this range on the meter
10nF | 0.05 | 0 | PASS | PASS | 0 | PASS | PASS
100nF | 0.5 | 0 | PASS | PASS | 0 | PASS | PASS
1uF | 0.005 | 0 | PASS | PASS | 0 | PASS | PASS
10uF | 0.05 | 0 | PASS | PASS | 0 | PASS | PASS
100uF | 0.5 | 0 | PASS | PASS | 0 | PASS | PASS
1mF | 0.005 | 0 | PASS | PASS | 0 | PASS | PASS
10mF | 0.05 | 0 | PASS | PASS | 0 | PASS | PASS
DC Voltage (Shorted terminals - 5 1/2 Digit mode)
100mV | 0.008 | 0 | PASS | PASS | 0 | PASS | PASS
1V | 5e-05 | 0 | PASS | PASS | 0 | PASS | PASS
10V | 0.0005 | 0 | PASS | PASS | 0 | PASS | PASS
100V | 0.005 | 0 | PASS | PASS | 0 | PASS | PASS
1kV | 0.05 | 0 | PASS | PASS | 0 | PASS | PASS
Resistance 2-Wire (Shorted terminals - 5 1/2 Digit mode)
100 Ohms | 0.208 | 0.001 | PASS | PASS | 0.001 | PASS | PASS
1k Ohms | 0.00028 | 1e-05 | PASS | PASS | 1e-05 | PASS | PASS
10k Ohms | 0.0007 | 0 | PASS | PASS | 0 | PASS | PASS
100k Ohms | 0.005 | 0 | PASS | PASS | 0 | PASS | PASS
1M Ohms | 5e-05 | 0 | PASS | PASS | 0 | PASS | PASS
10M Ohms | 0.0005 | 0 | PASS | PASS | 0 | PASS | PASS
100M Ohms | 0.005 | 0 | PASS | PASS | 0 | PASS | PASS
|  | As Found |  |  | As Left
Zero Offset | Tolerance | 34450A | 100% | 50% | 34450A | 100% | 50%
Resistance 4-Wire (Shorted terminals - 5 1/2 Digit mode)
100 Ohms | 0.008 | 0.001 | PASS | PASS | 0.002 | PASS | PASS
1k Ohms | 8e-05 | 1e-05 | PASS | PASS | 1e-05 | PASS | PASS
10k Ohms | 0.0005 | 0.0001 | PASS | PASS | 0 | PASS | PASS
100k Ohms | 0.005 | 0.001 | PASS | PASS | 0.001 | PASS | PASS
1M Ohms | 5e-05 | 0 | PASS | PASS | 0 | PASS | PASS
10M Ohms | 0.0005 | 0.0001 | PASS | PASS | 0.0001 | PASS | PASS
100M Ohms | 0.005 | 0 | PASS | PASS | 0 | PASS | PASS
|  | As Found |  |  | As Left
DC Voltage | Tolerance | 34450A | 100% | 50% | 34450A | 100% | 50%
100mV | 0.026 | 99.991 | PASS | PASS | 100 | PASS | PASS
-100mV | 0.026 | -99.991 | PASS | PASS | -100.001 | PASS | PASS
1V | 0.0002 | 0.99994 | PASS | PASS | 1 | PASS | PASS
-1V | 0.0002 | -0.99992 | PASS | PASS | -1 | PASS | PASS
10V | 0.002 | 9.9994 | PASS | PASS | 10 | PASS | PASS
100V | 0.02 | 99.988 | PASS | FAIL | 100 | PASS | PASS
1kV | 0.2 | 999.91 | PASS | PASS | 1000 | PASS | PASS
AC Voltage
100mV @20Hz | 1.1 | 99.936 | PASS | PASS | 99.93 | PASS | PASS
100mV @45Hz | 0.3 | 99.975 | PASS | PASS | 99.974 | PASS | PASS
100mV @1kHz | 0.3 | 99.988 | PASS | PASS | 99.99 | PASS | PASS
100mV @10kHz | 0.3 | 99.978 | PASS | PASS | 100.012 | PASS | PASS
100mV @30kHz | 1.8 | 99.878 | PASS | PASS | 99.958 | PASS | PASS
100mV @100kHz | 3.3 | 99.841 | PASS | PASS | 99.938 | PASS | PASS
1V @20Hz | 0.011 | 0.99931 | PASS | PASS | 0.99942 | PASS | PASS
1V @45Hz | 0.003 | 0.99972 | PASS | PASS | 0.99983 | PASS | PASS
1V @1kHz | 0.003 | 0.99984 | PASS | PASS | 0.99995 | PASS | PASS
1V @10kHz | 0.003 | 0.9997 | PASS | PASS | 0.99986 | PASS | PASS
1V @30kHz | 0.018 | 0.9992 | PASS | PASS | 0.99946 | PASS | PASS
1V @100kHz | 0.033 | 1.0026 | PASS | PASS | 1.003 | PASS | PASS
10V @ 20Hz | 0.11 | 9.9936 | PASS | PASS | 9.9949 | PASS | PASS
10V @ 45Hz | 0.03 | 9.9974 | PASS | PASS | 9.9992 | PASS | PASS
10V @ 1KkHz | 0.03 | 9.998 | PASS | PASS | 9.9998 | PASS | PASS
10V @ 10kHz | 0.03 | 9.9966 | PASS | PASS | 9.9995 | PASS | PASS
10V @ 30kHz | 0.18 | 9.9962 | PASS | PASS | 10.0006 | PASS | PASS
10V @ 100kHz | 0.33 | 10.0704 | PASS | PASS | 10.0752 | PASS | PASS
|  | As Found |  |  | As Left
AC Voltage | Tolerance | 34450A | 100% | 50% | 34450A | 100% | 50%
100V @45Hz | 0.3 | 99.981 | PASS | PASS | 99.99 | PASS | PASS
100V @1kHz | 0.3 | 99.99 | PASS | PASS | 99.999 | PASS | PASS
100V @10kHz | 0.3 | 99.995 | PASS | PASS | 100.019 | PASS | PASS
100V @30kHz | 1.8 | 99.85 | PASS | PASS | 99.89 | PASS | PASS
100V @100kHz | 3.3 | 98.342 | PASS | FAIL | 98.376 | PASS | PASS
750V @45Hz | 2.25 | 749.75 | PASS | PASS | 749.89 | PASS | PASS
750V @1kHz | 2.25 | 749.84 | PASS | PASS | 749.99 | PASS | PASS
750V @10kHz | 2.25 | 749.56 | PASS | PASS | 750.23 | PASS | PASS
Capacitance
1nF | No Specification, skip this range on the meter
10nF | 0.15 | 10.01 | PASS | PASS | 10.01 | PASS | PASS
100nF | 1.5 | 100 | PASS | PASS | 100 | PASS | PASS
1uF | 0.015 | 1 | PASS | PASS | 1 | PASS | PASS
10uF | 0.15 | 9.99 | PASS | PASS | 9.99 | PASS | PASS
100uF | 1.5 | 100 | PASS | PASS | 100 | PASS | PASS
1mF | 0.015 | 1 | PASS | PASS | 1 | PASS | PASS
10mF | 0.25 | 10 | PASS | PASS | 10 | PASS | PASS
Resistance 2-Wire
100 Ohms | 0.058 | 100.005 | PASS | PASS | 100.005 | PASS | PASS
1k Ohms | 0.00058 | 0.99995 | PASS | PASS | 0.99995 | PASS | PASS
10k Ohms | 0.0055 | 9.9996 | PASS | PASS | 9.9996 | PASS | PASS
100k Ohms | 0.055 | 99.994 | PASS | PASS | 99.994 | PASS | PASS
1M Ohms | 0.00065 | 0.99996 | PASS | PASS | 0.99996 | PASS | PASS
10M Ohms | 0.0255 | 9.9999 | PASS | PASS | 9.9999 | PASS | PASS
100M Ohms | 2.005 | 100.489 | PASS | PASS | 100.489 | PASS | PASS
Resistance 4-Wire
100 Ohms | 0.058 | 99.984 | PASS | PASS | 100.001 | PASS | PASS
1k Ohms | 0.00058 | 0.99989 | PASS | PASS | 1 | PASS | PASS
10k Ohms | 0.0055 | 9.999 | PASS | PASS | 10 | PASS | PASS
100k Ohms | 0.055 | 99.991 | PASS | PASS | 100 | PASS | PASS
1M Ohms | 0.00065 | 1.0001 | PASS | PASS | 1 | PASS | PASS
10M Ohms | 0.0255 | 10.0107 | PASS | PASS | 10.0007 | PASS | PASS
100M Ohms | 2.005 | 100.365 | PASS | PASS | 100.079 | PASS | PASS
|  | As Found |  |  | As Left
DC Current | Tolerance | 34450A | 100% | 50% | 34450A | 100% | 50%
100uA | 0.065 | 99.985 | PASS | PASS | 99.985 | PASS | PASS
1mA | 0.00057 | 0.99987 | PASS | PASS | 0.99987 | PASS | PASS
10mA | 0.0065 | 9.998 | PASS | PASS | 9.998 | PASS | PASS
100mA | 0.057 | 99.993 | PASS | PASS | 99.993 | PASS | PASS
1A | 0.00115 | 0.99948 | PASS | PASS | 0.99948 | PASS | PASS
10A | 0.0257 | 10.0043 | PASS | PASS | 10.0043 | PASS | PASS
AC Current
10mA @ 20Hz | 0.16 | 9.9913 | PASS | PASS | 9.9913 | PASS | PASS
10mA @ 45Hz | 0.06 | 9.9955 | PASS | PASS | 9.9955 | PASS | PASS
10mA @ 1kHz | 0.06 | 9.997 | PASS | PASS | 9.997 | PASS | PASS
10mA @ 5kHz | 0.22 | 10.0055 | PASS | PASS | 10.0055 | PASS | PASS
10mA @ 10kHz | 0.22 | 10.0195 | PASS | PASS | 10.0195 | PASS | PASS
100mA @ 20Hz | 1.6 | 99.931 | PASS | PASS | 99.931 | PASS | PASS
100mA @ 45Hz | 0.6 | 99.971 | PASS | PASS | 99.971 | PASS | PASS
100mA @ 1kHz | 0.6 | 99.989 | PASS | PASS | 99.989 | PASS | PASS
100mA @ 5kHz | 2.2 | 100.089 | PASS | PASS | 100.089 | PASS | PASS
100mA @ 10kHz | 2.2 | 100.267 | PASS | PASS | 100.267 | PASS | PASS
1A @ 20Hz | 0.016 | 0.99909 | PASS | PASS | 0.99909 | PASS | PASS
1A @ 45Hz | 0.006 | 0.99947 | PASS | PASS | 0.99947 | PASS | PASS
1A @ 1kHz | 0.006 | 0.99944 | PASS | PASS | 0.99944 | PASS | PASS
1A @ 5kHz | 0.022 | 1.0012 | PASS | PASS | 1.0012 | PASS | PASS
2A @ 5kHz | 0.06 | 2.0003 | PASS | PASS | 2.0003 | PASS | PASS
10A @ 45Hz | 0.06 | 10.0031 | PASS | PASS | 10.0031 | PASS | PASS
10A @ 1kHz | 0.06 | 10.0033 | PASS | PASS | 10.0033 | PASS | PASS
Frequency
20Hz @ .1V | 0.007 | 20 | PASS | PASS | 20 | PASS | PASS
1kHz @ 1V | 0.00023 | 1 | PASS | PASS | 1 | PASS | PASS